## Check Sheet
Rev. | [5 empty] | New York State Department Of Environmental Conservation | [25 empty] | Reviewed By:
2014-09-01
TRACE LEVEL CARBON MONOXIDE (CO) ANALYZER
[6 empty] | Weekly Check Sheet | [25 empty] | Date:
Station: |  |  | Pinnacle State Park | [10 empty] | Station #: |  |  |  | 500104PRI | [8 empty] | Week Of: |  |  |  |  | 2020-01-05
Operator: |  |  | John B Spicer | [9 empty] | Performed By: |  |  |  |  | John B Spicer | [6 empty] | Day of Week: | [6 empty] | Tuesday
|  |  | Date of Testing: | [6 empty] | 2020-01-07 | [5 empty] | Time: |  |  | OOC: |  | 10:04 |  |  |  | IC: | 11:52 |  |  | hours EST
[11 empty] | Analyzer Data | [11 empty] | ** if Audit or Calibration dates unknown leave blank.
Model #: |  |  |  | API 300EU | [6 empty] | ENCON #: | [5 empty] | 159662 | [6 empty] | Serial #: | [7 empty] | 132
| Last 2 Calibration Dates: | [8 empty] | 2018-11-13 |  |  |  | 2019-03-29 |  |  |  |  | Last 2 Audit Dates: | [7 empty] | 2019-09-10 | [5 empty] | 2019-11-06
[11 empty] | Calibrator Information | [11 empty] | ** taken from calibrator QA certification sheet
Calibrator #: | [6 empty] | T700U |  |  |  | Cert. Period: |  |  |  |  | 2019-08-21 |  |  |  | ↔ | 2020-02-28 |  |  |  |  | CO Tank Conc., ppm: | [8 empty] | 300
CO Primary Tank: | [6 empty] | CC408646 |  |  |  | Cylinder Cert. Expires: | [7 empty] | 2020-02-28 | [6 empty] | Cylinder Press., psig: | [8 empty] | 900
Zero, Span and PC Checks
[5 empty] | % of |  |  |  | QA Calibrator | [12 empty] | Data Logger
[5 empty] | Full |  |  |  | Dilution |  |  | Pollutant |  |  |  |  | Certified |  |  |  |  | Measured | [5 empty] | % |  |  |  | Date
** If necessary |  |  |  |  | Scale |  |  |  | Setting |  |  | Setting |  |  |  |  | Conc., ppb |  |  |  |  | Conc., ppb | [5 empty] | Diff ‡ |  |  |  | Completed
Zero start** |  |  |  |  | 0% |  |  |  | 8 |  |  | Off |  |  |  |  | 0 |  |  |  |  | 20 | [5 empty] | -------- |  |  |  | 2020-01-07
Span |  |  |  |  | 95% |  |  |  | 8 | [7 empty] | 3509 |  |  |  |  | 3593 | [5 empty] | 2.4 |  |  |  | 2020-01-07
PC |  |  |  |  | 25% |  |  |  | 8 | [7 empty] | 938 |  |  |  |  | 927 | [5 empty] | -1.2 |  |  |  | 2020-01-07
Zero end |  |  |  |  | 0% |  |  |  | 8 |  |  | Off |  |  |  |  | 0 |  |  |  |  | 20 |  |  | -30 |  |  | -------- |  |  |  | 2020-01-07
CO Offset: |  |  |  |  | as found→ | [6 empty] | as left → |  |  | ------ |  |  | note: if no adjustment is made to | [9 empty] | 0.6 |  |  |  | <= Avg. % Diff
CO Offset leave "as left" blank.
‡ { Span Limits → Warning: ± 3 %, Control: ± 6 % / Zero Drift → ± 30 ppb / PC Limits → Warning: ± 4 %, Control: ± 8 %}
Scheduled Maintenance
Procedure | [11 empty] | Frequency | [8 empty] | Last Done | [8 empty] | BY
Replace Particulate Filter | [11 empty] | As Needed ** | [8 empty] | 2019-12-31 | [8 empty] | JBS
Flow Check | [11 empty] | Annually | [8 empty] | 2019-12-31 | [8 empty] | JBS
Leak Check | [11 empty] | Annually
Replace Pump Diaphragm | [11 empty] | Annually
** Performed more often as required by local conditions.
[9 empty] | See other side for Diagnostics and any Comments. | [24 empty] | Page 1 of 2
Diagnostic Checks
Range |  |  |  | 2000 ppb |  |  |  | 1000 |  |  | Sample |  |  |  |  | 28.4 ± 1" |  |  | 26.8 |  |  |  |  | CO |  |  |  |  | Last |  |  |  | 1.092
[11 empty] | Pressure | [12 empty] | Slope |  |  |  |  | Calibration
STABIL |  |  |  | < 10 ppb |  |  |  | 947 |  |  | Sample |  |  |  |  | 1.8 |  |  | 1803 |  |  |  |  | CO |  |  |  |  | Last |  |  |  | -0.02048
|  |  |  | w/ zero air | [6 empty] | Flow |  |  |  |  | ± .36 lpm | [7 empty] | Offset |  |  |  |  | Adjustment
CO MEAS |  |  |  | 4300 - 4700 |  |  |  | 4052 |  |  | Bench |  |  |  |  | 48° |  |  | 48 |  |  |  |  | Shelter Temp. | [6 empty] | deg C |  | 21
|  |  |  | w/ zero air | [6 empty] | Temp. |  |  |  |  | ± 1° C
CO REF |  |  |  | 3600 - 4000 |  |  |  | 3432 |  |  | Wheel |  |  |  |  | 62° |  |  | 62 |  |  |  |  | Shelter BP | [6 empty] | inHg |  | 28.1
|  |  |  | w/ zero air | [6 empty] | Temp. |  |  |  |  | ± 1° C
MR Ratio |  |  |  | 1.18 ± .001 |  |  |  | 1.18 |  |  | Oven |  |  |  |  | 48° |  |  | 46.8
|  |  |  | w/ zero air | [6 empty] | Temp. |  |  |  |  | ± 1° C
AZERO |  |  |  | 1.165 - |  |  |  | 1.187 |  |  | PHT |  |  |  |  | 250 - |  |  | 2585
Ratio |  |  |  | 1.225 | [6 empty] | Drive |  |  |  |  | 4750 mv
Are analyzer diagnostics within tolerances? | [12 empty] | Y / N |  |  | y |  |  | If no please explain in comments section.
Comments:
Page 2 of 2
Formulas:
Date = IF(AND(ISNUMBER(L11),ISNUMBER([Certified]),ISNUMBER([Measured])),L11,"-------- ")
Date = IF(AND(ISNUMBER(L11),ISNUMBER([Certified]),OR(ISNUMBER([Measured]),ISNUMBER(AD42))),L11,"-------- ")
Note on A1: Red comment flags provide useful information throughout form.
Note on AK7: Date entered automatically. Note: week defined as starting on Sunday and ending on Saturday.
Note on L11: note: can press CNTRL ; simultaneously to enter current date in a cell.
Note on A18: Completed minimum once per year. See QA manual section 7.3.0 for other instances
Note on W18: Completed minimum once per quarter. See QA manual sections 7.1.7 & 7.2.0 for other instances.
Note on A36: Completed minimum once per week.
Note on AK36: For Zero/Span cells: date entered automatically. If cell remains blank verify Date of Testing cell and Conc. ppb cells have values.
Note on A38: Completed minimum once per week.
Note on A40: Completed minimum once every 2 weeks.
Note on AK40: Enter date manually. Note: can press CNTRL ; simutaneously to enter current date in this cell.
Note on AL74: Enter Value in Celsius or Fahrenheit.
Note on AL76: Enter value in mmHg or inHg.

## Limit Verification
Verification of QA Limit Exeedances
| Span Warning Limit Exceedance: (Yes / No) → | No
| Span Control Limit Exceedance: (Yes / No) → | No
| Start Zero Drift Exceedance: (Yes / No) → | No
| End Zero Drift Exceedance: (Yes / No) → | No
| PC Warning Limit Exceedance: (Yes / No) → | No
| PC Control Limit Exceedance: (Yes / No) → | No
**Note: Refer to QA manual for guidance in the event of an exceedance.