## Check Sheet
Rev. | [5 empty] | New York State Department Of Environmental Conservation | [25 empty] | Reviewed By:
2014-09-01
43i TRACE LEVEL SULFUR DIOXIDE (SO2) ANALYZER
[6 empty] | Weekly Check Sheet | [25 empty] | Date:
Station: |  |  | Pinnacle State Park | [10 empty] | Station #: |  |  |  | 500104PRI | [8 empty] | Week Of: |  |  |  |  | 2019-12-29
Operator: |  |  | John B Spicer | [9 empty] | Performed By: |  |  |  |  | John B Spicer | [6 empty] | Day of Week: | [6 empty] | Thursday
|  |  | Date of Testing: | [5 empty] | 2020-01-02 | [6 empty] | Time: |  |  | OOC: |  | 10:18 |  |  |  | IC: | 11:15 |  |  | hours EST
[11 empty] | Analyzer Data | [11 empty] | ** if Audit or Calibration dates unknown leave blank.
Model: |  |  |  |  | 43i-TLE | [5 empty] | Encon #: | [5 empty] | 155351 | [6 empty] | Serial #: |  |  |  |  | 617916895
| Last 2 Calibration Dates: | [7 empty] | 2018-03-13 |  |  |  |  | 2019-01-15 |  |  |  |  | Last 2 Audit Dates: | [7 empty] | 2019-09-09 | [5 empty] | 2019-11-05
[11 empty] | Calibrator Information | [11 empty] | ** taken from calibrator certification sheet
Model: |  |  |  |  | 700E | [5 empty] | Calibrator #: | [5 empty] | 292 | [7 empty] | Cylinder Number: | [7 empty] | CAL3633
Calibrator Certification Period: | [10 empty] | 2019-08-21 | [5 empty] | thru | 2020-02-28 | [6 empty] | Cert. Conc., (ppm) | [8 empty] | 9.9
Zero, Span and PC Checks
[5 empty] | % of |  |  | QA Calibrator | [13 empty] | Data Logger
[5 empty] | Full |  |  | Dilution |  |  |  | Pollutant |  |  |  |  | Certified |  |  |  |  | Measured | [5 empty] | % |  |  |  | Date
[5 empty] | Scale |  |  | Setting |  |  |  | Setting |  |  |  |  | Conc., ppb |  |  |  |  | Conc., ppb | [5 empty] | Diff ‡‡ |  |  |  | Completed
Span |  |  |  |  | 90% |  |  | 5 | [8 empty] | 91.7 |  |  |  |  | 90.1 | [5 empty] | -1.7 |  |  |  | 2020-01-02
PC |  |  |  |  | 5% |  |  | 10 | [8 empty] | 5.2 |  |  |  |  | 5 | [5 empty] | -3.8 |  |  |  | 2020-01-02
Zero ** |  |  |  |  | 0% |  |  | 5 |  |  |  | Off |  |  |  |  | 0 |  |  |  |  | -0.1 |  |  | -0.1 |  |  | ------- |  |  |  | 2020-01-02
** SO2 Background: | [5 empty] | as found→ |  |  | 2.3 |  |  | as left → |  |  | ----- |  |  | note: if no adjustment is made to | [9 empty] | -2.75 |  |  |  | <= Avg. % Diff
bkgd factor leave "as left" blank.
‡‡ { Span Limits → Warning: ± 5 %, Control: ± 9 % / Zero Drift → ± 2 ppb / PC Limits → Warning: ± 6 %, Control: ± 10 %}
Scheduled Maintenance
Procedure | [11 empty] | Frequency | [6 empty] | Last Done | [6 empty] | Next Due | [6 empty] | BY
Replace Particulate Filter | [11 empty] | Monthly | [6 empty] | 2019-12-31 | [6 empty] | 2020-01-31 | [6 empty] | JBS
Check for leaks | [11 empty] | Monthly
Check/Clean Capillary | [11 empty] | As Needed
Clean Ventilation Fan Screen | [11 empty] | Monthly | [6 empty] | 2019-12-31 | [6 empty] | 2020-01-31 | [6 empty] | JBS
Flash Lamp Replaced | [11 empty] | As Needed | [6 empty] | 2015-03-26
[7 empty] | See other side for Diagnostic Checks and Comments. | [26 empty] | Page 1 of 2
Diagnostic Checks
Range |  |  |  | (ppb) |  | 100 |  |  |  | Voltages | [5 empty] | Motherboard |  |  |  | Interface
Avg. Time |  |  |  | (secs) |  | 60 |  |  |  | PMT Volts | [9 empty] | -632.3 |  |  |  |  | Internal Temp | [5 empty] | (° C) |  | 31.6
SO2 Bkgrnd |  |  | Last |  |  | 2.87 |  |  |  | Flash Volts | [9 empty] | 1109 |  |  |  |  | Chamber Temp | [5 empty] | (° C) |  | 44.9
Calibration
SO2 Coeff. |  |  | Last |  |  | 1.022 |  |  |  | Supply |  |  | (+ 3.3V) |  |  | 3.3 |  |  |  | 3.2 | [5 empty] | Pressure |  |  |  | (mmHg) |  |  | 695.5
Calibration
Alarm On? | [5 empty] | n |  |  |  | Supply |  |  | (+ 5V) |  |  | 5 |  |  |  | 5 |  |  |  |  | Flow | [5 empty] | (LPM) |  | 0.464
Temp Corr. On? | [5 empty] | y |  |  |  | Supply |  |  | (+15V) |  |  | 15.1 |  |  |  | 14.6 |  |  |  |  | Lamp Intensity | [5 empty] | (%) |  | 89
Press. Corr. On? | [5 empty] | y |  |  |  | Supply |  |  | (- 15V) | [6 empty] | -15.1 |  |  |  |  | Shelter Temp. |  |  |  |  | (deg C) |  |  | 19.5
Flash Lamp On? | [5 empty] | y |  |  |  | Supply |  |  | (24V) |  |  | 24.1 |  |  |  | 23.8 |  |  |  |  | Shelter BP |  |  |  |  | (inHg) |  |  | 27.96
Service Off? | [5 empty] | y |  |  |  | Supply |  |  | (- 3.3V) |  |  | -3.2
Are analyzer diagnostics within tolerances? | [12 empty] | Y / N |  |  | y |  |  | If no please explain in comments section.
Comments:
Page 2 of 2
Formulas:
Date = IF(ISBLANK(L11)," ",IF(ISNUMBER([Diff ‡‡]),L11,"-------- "))
Date = IF(ISBLANK(L11)," ",IF(AND(ISNUMBER([Certified]),OR(ISNUMBER(AD40),ISNUMBER([Measured]))),L11,"-------- "))
Note on A1: Red comment flags provide useful information throughout form.
Note on AK7: Date entered automatically. Note: week defined as starting on Sunday and ending on Saturday.
Note on L11: note: can press CNTRL ; simultaneously to enter current date in a cell.
Note on A18: Completed minimum once per year. See QA manual section 5.1.4 for other instances.
Note on W18: Completed minimum once per quarter. See QA manual sections 5.1.4 & 5.1.6 for other instances.
Note on A36: Completed minimum once per week. Reference: QA manual section 5.1.4
Note on AK36: For Zero/Span cells: date entered automatically. If cell remains blank verify Date of Testing cell and chart Data cells have values in them.
Note on A38: Completed minimum once every 2 weeks. Reference: QA manual section 5.1.4
Note on AK38: Enter date manually. Note: can press CNTRL + ; simultaneously to enter current date in this cell.
Note on A40: Completed minimum once per week.
Note on AL80: Enter Value in Celsius or Fahrenheit.
Note on AL82: Enter value in mmHg or inHg.

## Limit Verification
Verification of QA Limit Exeedances
| Span Warning Limit Exceedance: (Yes / No) → | No
| Span Control Limit Exceedance: (Yes / No) → | No
| Zero Drift Exceedance: (Yes / No) → | No
| PC Warning Limit Exceedance: (Yes / No) → | No
| PC Control Limit Exceedance: (Yes / No) → | No
**Note: Refer to QA manual for guidance in the event of an exceedance.