## Check Sheet
Rev. | [5 empty] | New York State Department Of Environmental Conservation | [23 empty] | Reviewed By:
2014-09-01 | [5 empty] | 49i OZONE (O3) ANALYZER
[6 empty] | Weekly Check Sheet | [23 empty] | Date:
Station: |  |  | Pinnacle State Park | [9 empty] | Station #: |  |  | 500104PRI | [8 empty] | Week Of: |  |  |  |  | 2020-01-05
Operator: |  |  | John B Spicer | [8 empty] | Performed By: |  |  |  | John B Spicer | [6 empty] | Day of Week: | [6 empty] | Tuesday
|  |  | Date of Testing: |  |  |  |  | 2020-01-07 | [5 empty] | Time: |  |  | OOC: |  | 12:00 |  |  |  | IC: | 14:10 |  |  | hours EST
[10 empty] | Analyzer Data | [10 empty] | ** if Audit or Calibration dates unknown leave blank.
Model: |  |  |  |  | TEI 49i |  |  |  |  | Encon #: |  |  |  |  | 155356 | [6 empty] | Serial #: |  |  |  |  | 617916613
| Last 2 Calibration Dates: | [6 empty] | 2018-08-16 |  |  |  |  | 2019-07-02 |  |  |  | Last 2 Audit Dates: | [7 empty] | 2019-07-02 | [5 empty] | 2019-11-05
[10 empty] | Calibrator Information | [10 empty] | ** taken from calibrator QA certification sheet
Model: |  |  |  |  | API T700 |  |  |  |  | Encon #: |  |  |  |  | 177073 | [7 empty] | Lab BP → | [5 empty] | 29.82 |  |  |  | mmHg
Calibrator Certification Period: | [9 empty] | 2019-08-21 |  |  |  |  | thru | 2020-02-28
Zero, Span and PC Checks
[5 empty] | % of |  | QA Calibrator | [12 empty] | Data Logger
[5 empty] | Full |  | Dilution |  |  |  | Pollutant |  |  |  | Certified |  |  |  |  | Measured | [5 empty] | % |  |  |  | Date
[5 empty] | Scale |  | Setting |  |  |  | Setting |  |  |  | Conc., ppm |  |  |  |  | Conc., ppm | [5 empty] | Diff ‡‡ |  |  |  | Completed
Span |  |  |  |  | 85% |  | 6 | [7 empty] | 0.1256 |  |  |  |  | 0.1244 | [5 empty] | -1 |  |  |  | 2020-01-07
PC |  |  |  |  | 15% |  | 6 | [7 empty] | 0.0651 |  |  |  |  | 0.0644 | [5 empty] | -1.1 |  |  |  | 2020-01-07
Zero ** |  |  |  |  | 0% |  | 6 |  |  |  | Off |  |  |  | 0 |  |  |  |  | 0.0002 |  |  | 0 |  |  | -------- |  |  |  | 2020-01-07
** O3 Background: | [5 empty] | as found→ |  |  |  |  | as left → |  |  |  |  | note: if no adjustment is made to | [9 empty] | -1.05 |  |  |  | <= Avg. % Diff
bkgd factor leave "as left" blank.
‡‡ { Span Limits → Warning: ± 5 %, Control: ± 9 % / Zero Drift → ± .003 ppm / PC Limits → Warning: ± 6 %, Control: ± 10 %}
Scheduled Maintenance
Procedure | [10 empty] | Frequency | [5 empty] | Last Done | [6 empty] | Next Due | [6 empty] | BY
Replace Particulate Filter | [10 empty] | As Needed | [5 empty] | 2019-12-31 | [6 empty] | 2020-01-31 | [6 empty] | JBS
Check / Clean Cooling Fan Filter | [10 empty] | Monthly | [5 empty] | 2019-12-31 | [6 empty] | 2020-01-31 | [6 empty] | JBS
Check for leaks in solenoids | [10 empty] | Every Audit/Cal
Test Ozone Scrubber Efficiency | [10 empty] | Every Audit/Cal
Check or Clean Sample Cells | [10 empty] | As Needed
[7 empty] | See other side for Diagnostics and Comments. | [24 empty] | Page 1 of 2
Diagnostics
Range |  |  |  | (ppm) |  | 0.5 |  |  | Voltages |  |  |  |  | Motherboard |  |  |  | Interface |  |  |  |  | Bench Temp. | [5 empty] | (°C) |  | 33.3
Avg. Time |  |  |  | (secs) |  | 60 |  |  | Supply |  |  | + 3.3V |  | 3.3 |  |  |  | 3.3 |  |  |  |  | Bench Lamp | [5 empty] | (°C) |  | 53.5
Temp.
O3 Bkgrnd |  |  | Last |  |  | 0.0008 |  |  | Supply |  |  | + 5.0V |  | 5 |  |  |  | 5 |  |  |  |  | O3 Lamp Temp. | [5 empty] | (°C) |  | 67.5
Calibration
O3 Coeff. |  |  | Last |  |  | 1.001 |  |  | Supply |  |  | +15.0V |  | 14.9 |  |  |  | 15.1 |  |  |  |  | Pressure |  |  |  |  | (mmHg) |  |  | 706.2
Calibration
Alarm On? | [5 empty] | n |  |  | Supply |  |  | - 15.0V | [5 empty] | -15.1 |  |  |  |  | Flow |  |  |  |  | Cell A |  |  | 0.733
Temp Corr. On? | [5 empty] | y |  |  | Supply |  |  | 24.0V |  | 23.9 |  |  |  | 23.7 |  |  |  |  | (0.4 - 1.4 |  |  |  |  | Cell B |  |  | 0.683
lpm)
Press. Corr. On? | [5 empty] | y |  |  | Supply |  |  | - 3.3V |  | -3.2 | [8 empty] | Intensities |  |  |  |  | Cell A |  |  | 53397
O3 Solenoid Off? | [5 empty] | y |  |  | Photo Lamp | [8 empty] | 10 |  |  |  |  | (>45Khz - |  |  |  |  | Cell B |  |  | 70534
< 150 Khz)
Service Off? | [5 empty] | y |  |  | O3 Lamp | [8 empty] | 5.8 |  |  |  |  | Lamp Setting | [5 empty] | (%)
Shelter Temp |  |  |  | deg C |  | 19.9 | [16 empty] | Reference |  |  |  |  | Cell A
Shelter BP |  |  |  | inHg |  | 28.06 | [16 empty] | Noise |  |  |  |  | Cell B
(< 4.0)
Are analyzer diagnostics within tolerances? | [11 empty] | Y / N |  | y |  |  | If no please explain in comments section.
Comments:
Page 2 of 2
Formulas:
Date = IF(OR(ISBLANK(L10),ISBLANK([Diff ‡‡])),"-------- ",IF(ISNUMBER([Diff ‡‡]),L10,"--------"))
Date = IF(ISBLANK(L10),"--------",IF(AND(ISNUMBER([Certified]),OR(ISNUMBER([Measured]),ISNUMBER([Lab BP →]))),L10,"--------"))
Note on AK6: Date entered automatically. Note: week defined as starting on Sunday and ending on Saturday.
Note on L10: note: can press CNTRL + ; simultaneously to enter current date in a cell.
Note on A17: Completed minimum once per year. See QA manual section 6.3.0 for other instances
Note on W17: Completed minimum once per quarter. See QA manual sections 6.1.7 & 6.2.0 for other instances.
Note on A35: Completed minimum once per week. Reference: QA manual section 6.1.4
Note on AK35: For Zero/Span: date entered automatically. If cell remains blank verify Date of Testing and chart Conc., ppm cells have values.
Note on A37: Completed minimum once every 2 weeks. Reference: QA manual section 6.4.0
Note on AK37: For PC: Enter date manually. Can press CNTRL + ; simultaneously to enter current date in this cell.
Note on A39: Completed minimum once per week.

## Limit Verification
Verification of QA Limit Exeedances
| Span Warning Limit Exceedance: (Yes / No) → | No
| Span Control Limit Exceedance: (Yes / No) → | No
| Zero Drift Exceedance: (Yes / No) → | No
| PC Warning Limit Exceedance: (Yes / No) → | No
| PC Control Limit Exceedance: (Yes / No) → | No
**Note: Refer to QA manual for guidance in the event of an exceedance.